Field experiment preparation work for Steve Edburg
2012 June 15
Patrick O'Keeffe / patrick.okeeffe@email.wsu.edu / 509.335.7246

Equipment
- The two ATI sonics (SN 410 & 605) have been configured to work properly with the data logger program.
- There was a ~3 ºC difference in temperature output between the sonics. Their last calibration is unknown.
- To connect to the data logger using PC200, you will need a computer with a serial port or a usb-serial adapter. Since all the data is stored on the card though, there is not much reason to connect to the logger.
- To access the data on the card you will need a compact flash reader. Be sure to press the white safe removal button or power down the logger before removing the card.

Data logger program
- The data logger is loaded with the ATI_recorder.cr1 program, which is set to run on power-up. The compact flash card was formatted after testing.
- Only the 10hz U, V, W, Ts, and diagnostic word are logged.
- Each sonic is logged in its own file
- The tables permit 14 days of data logging, although that can be changed by modifying line 20 of the program file using any plain-text editor. Using the total amount of space available, the program can log data for ~60 days. Choose a sufficiently long duration but avoid storing empty days to keep file sizes reasonable.

Diagnostic word
The last column in the data files is a diagnostic code which summarizes the various 'diagnostics' provided by the sonic anemometer. It consists of an integer in the range of 0 to 2222.
Sample diagnotic word:	1	0	1	0	1 = axis blockage / transducer failure
Corresponding signal:	Ts	W	V	U	2 = data quality algorithm rejection

Sample header for TOA5 data file
"TOA5","LOCAL_31","CR1000","34375","CR1000.Std.22","CPU:ATI_recorder.cr1","30799","sonic_com1"
"TIMESTAMP","RECORD","Ux_1","Uy_1","Uz_1","Ts_1","diag_1"
"TS","RN","m/s","m/s","m/s","C","m/s"
"","","Smp","Smp","Smp","Smp","Smp"

Table 1: ATI sonic configuration
Data mode: ASCII / binary	Data quality algorithm: ON
ASCII mode: terse / verbose	Shadow correction: ON
Baud rate: 9600	Remove RH from temp calc: OFF
Parity: even	Output: U,V,W,Ts at 10hz
data bits: 7	Data pipline delay: 0.1 seconds (not changeable)
stop bits: 1